TYPICAL SPECIFICATION

SOLO® G2

w/ SPILLSAFE™ DRUM SCALE

SPECIFICATION SECTION 11245 - Chemical Feeders

SPECIFICATION SECTION 11240 - Chemical Metering Pumps

A quantity of ____drum scales of___________ lbs (kg.) capacity shall be provided and shall be of the digital readout/electronic load cell type.  Scale platform shall be sized to accept up to a 26" (66cm) diameter drum.  Platform shall have integral secondary containment with a total capacity of 66 gallons, 11 gallons in the containment skid itself, and an additional 55 gallons in a roll out bag.  Overall platform height shall be no more than 7" (1780mm) to permit easy loading and unloading of drums.    Platform and containment system shall be non-metallic to eliminate the possibility of corrosion.

Scale shall be of the multiple load cell design.  Weight shall be transferred via a pivoted platform to a three (3) NTEP approved load cells of the shear beam strain gauge type.  Flexible cable shall connect scale platform to indicator to allow easy remote installation of the readout.  Cable length shall be _____ feet (20'/6M standard).

Indicator shall monitor__________(1 or 2 ) channels. The remote mounted LCD indicator shall carry CE marking and shall be housed in a NEMA 4X, UL approved enclosure. All operations shall be keypad operated & menu driven in order to avoid compromising the NEMA 4X seal at anytime.  The alphanumeric LCD readout shall have backlighting for readability in low light conditions.  Power requirement shall be ______________( 110 VAC, 220 VAC or 24 VDC)

A 6 digit numerical display shall give operator the ability to monitor chemical by weight (lb or kg) or volume (gallons or liters).  A bar graph display shall read 0-100% for the net contents.  A dual mode TARE key shall allow user to enter the tare weight of the vessel or enter the net weight of the chemical depending on application needs. A diagnostics menu shall allow recalibration without the need to apply field test weights.  A user adjustable filter function shall stabilize display in the event of vibration from pumps or mixers in the immediate vicinity of the scale.

Indicator shall output net weight via a 4-20mA signal and full scale output shall be user adjustable via the keypad.  Indicator shall have four adjustable set points to display low or high level conditions on the indicator.

OPTIONAL:    Set points shall activate ______ (up to four) Form C dry contact relays rated at 				2A@30VDC, 60W max power, 220VDC max voltage.

Scale shall carry a Full Five (5) Year Factory Warranty.  Limited warranties shall be considered unacceptable.   Full scale accuracy shall be better than 1/4 of 1%.  Scale shall be SpillSafe™ Model 26-SC6-HA3 with SOLO® G2 digital display,  Model SRG2-___ (1 or 2 channel), as manufactured by FORCE FLOW, 2430 Stanwell Drive, Concord, CA 94520 USA  (www.forceflow.com).